## Slide 1
Avalanche Photodiodes (APDs) testing at ESA/ESTEC
Vincent Affatato, Dennis Breeveld, Sander Blommaert, Luc Boucher, Frédéric Lemmel, Jean Le Graet, Eloy Hernandez, Malik Mansour, Thibaut Prod’homme, Conor Robinson, Erik Ryman, Peter Verhoeve
11/03/2026

## Slide 2: SCI-FIV Lab: Characterized FPAs in the last years
IBEX
ESA SCI-FIV Payload Technology Validation  Support for Future Science Missions

## Slide 3: Technology background
Saphira
320 x 256 or 512 x 512 pixels
Used as wavefront sensor for the VLT on Mount Paranal in Chile.
Ike Pono
1k x 1k or 2k x 2k
Developed in collaboration with University of Hawai’i and NASA. First large format LmAPD array optimised for good performance with glow and read noise.
IBEX
2k x 2k
Developed in collaboration with ESA. Large format array optimised for dark current performance.
GOAL:
Develop a large format NIR detector for ultra-low background applications (e.g. direct exoplanet detection) enabling single photon detection
BONUS:
Increase European technological capabilities in this sector

## Slide 4: The IBEX detector
Nominal temperature: 80 K
Geometrical features
2048 x 2048 pixels
16 / 8 / 4 channels
8 reference rows
Readout / Reset capabilities
Read reset read
Interleaved
Pixel by pixel, Line by line
Windowed
Timing parameters
Clock sampling at 270 kHz
Nominal detector readout in 1 s
Performance requirements
Dark current:
< 3 e-/p/s (at gain = 10)
Readout noise:
< 10 e- rms
CHC:
>100k e-
Glow:
?
Avalanche gain:
1 – 40
Cut-off:
> 2.5

## Slide 5: Cryostat Configurations
Simplified block diagram:
Cryo-Illumination
2
2
LAPD
86K
2 Flex
ARCHON
Cables
Cryo LEDs
200K
FWA - ND and
bandpass 80K
CPE
1
1
4
4
3
3
LAPD
86K
2 Flex
ARCHON
Cables
Cryo LEDs
200K
FWA - ND and
bandpass 80K
CPE
Aktar Cap
Simplified block diagram:
Dark Configuration
R2-D2

## Slide 6: Functionality characterization - Overview
Characterized Devices
Cryo-Illumination  PTC and Avalanche Gain
Dark configuration  Dark current, RON, Glow

## Slide 7: Functionality characterization - PTC
HW | Median log-log slope | CCF [e-/adu] | Sense node capacitance [fF] | FW var [ke-]
1 | 0.524 | 12.1 | 27.3 | 225
3 | 0.546 | 10.6 | 24.0 | 215
4 | 0.512 | 13.6 | 30.6 | 260
5 | 0.503 | 12.9 | 29.0 | 255
Frame subtraction across the ramps, evaluation of residual shot noise
Differential temporal approach
Estimated by LUK: 27.5

## Slide 8: Functionality characterization – Linearity and CHC
HW | CCF [e-/adu] | FW var [ke-] | FW linearity [ke-]
1 | 12.1 | 225 | 260
3 | 10.6 | 215 | 220
4 | 13.6 | 260 | 290
5 | 12.9 | 255 | 295
CHC >> 100 ke- from requirement
[adu]

## Slide 9: Functionality characterization – Avalanche gain
Power law
fit
b = 3.85, 3.98 Ike Pono = 3.01
data
Exp. law
b = 0.253, 0.225 Ike Pono = 0.255
LUK = 0.3
data
Avalanche gain << requirement (40)
Detector or controller?
y_bin

## Slide 10: Functionality characterization – Readout Noise
Left – right bias
Fowler approach
[adu]
Pair subtraction up the ramp, statistical evaluation based on each pair in different ramps
CDS readout noise
Readout noise higher than 10 e- requirement
E.g. HW1: 1 e- RON with 14 ke- FW

## Slide 11: Functionality characterization – Dark Signal (1)
4h acquisitions, continuously running but sampling only one frame per minute
Dark ramps
Self-generated glow, back reflected on the cap surface
Dark current seems to be below 2 e- at gain = 10

## Slide 12: Functionality characterization – Dark Signal (2)
Missing measurements at higher temperatures to evaluate physical phenomena causing dark signal generation
Plateau?

## Slide 13: Functionality characterization – Glow
Frame Rate
1 Hz
2 Hz
4 Hz
8 Hz
0.5, 0.25, 0.1 Hz
2h acquisitions, continuously running but sampling only one frame per minute at different SF rates
Glow ramps
Full frame, but after the clocking a buffer time of 1, 4, or 9 seconds is implemented
HW | Intercept | Slope [e-/frame] | 
1 | 0.024 | 0.23 | 0.998
3 | -0.026 | 0.38 | 0.951
4 | 0.078 | 0.26 | 0.993
5 | 0.035 | 0.23 | 0.998
HW | Intercept | Slope [e-/frame] | 
1 | 0.024 | 0.23 | 0.998
3 | -0.026 | 0.38 | 0.951
4 | 0.078 | 0.26 | 0.993
5 | 0.035 | 0.23 | 0.998
Glow contribution
Theoretical dark current
Glow seems to be the dominant contribution for dark signal

## Slide 14: Functionality characterization – Dark Signal (2) continued
Missing measurements at higher temperatures to evaluate physical phenomena causing dark signal generation
Correcting for the glow at low T
Applying Rule22 fit [Zandian 2023], we can estimate the cut-off wavelength ()
Cut-off predicted @ 3.25 , higher than requirement
Plateau?

## Slide 15: Functionality characterization – Summary
Performance results
Dark current:
< 2 e-/p/s (at gain = 10)
Below expected values
Readout noise:
< 15 e- rms
CHC:
>200k e-
Glow:
0.23 e-/p/s
Main dark signal contribution
Avalanche gain:
1 – 16
Cut-off:
e
Left – right bias
Fowler approach
High T?
Breakdown?
Controller?
Reference pixels correction?
Performance requirements
Dark current:
< 3 e-/p/s (at gain = 10)
Readout noise:
< 10 e- rms
CHC:
>100k e-
Glow:
?
Avalanche gain:
1 – 40
Cut-off:
> 2.5

## Slide 16: Next Steps - Collaborations
LAPD community
Technical Officer – Characterization Institute
Characterization Institute
–
Characterizing HW1 to get prepared for IBEX follow-up and on-sky demonstration
Characterization Institute –
Characterizing HW5 and preparing proton campaign
Max Planck Institute
–
Testing compatibility between IBEX and ALFA-c ASIC for the ASIDES mission
Manufacturer
–
Delivering IBEX follow-up this year

## Slide 17: Next steps - Tests
Inter Pixel Capacitance
Persistence
Quantum Efficiency
Cross Talk
Excess Noise
PRNU
RTS
Temperature sensitivity
Gain stability
Radiation sensitivity
Increase the TRL to make LmAPD a key technology for future European scientific missions